## Slide 1: Peer Stakeholder-Product Validation Review (PS-PVR) For the Provisional Maturity of GOES-18 SEISS MPS-LO L1b Product
November 29, 2022
GOES-R Calibration Working Group (CWG)
Trevor Leonard
CIRES CU Boulder / NOAA NCEI
Contributors: Athanasios Boudouridis, Brian Kress, Juan Rodriguez
CIRES CU Boulder / NOAA NCEI
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 2: Outline
MPS-LO Overview
Review of Beta/PLT Maturity
Product Quality Evaluation
  Instrument and GPA Anomalies and Resolution
  Analysis and PLPTs Supporting Assessment for Provisional Status
Provisional Maturity Assessment
Path to Full Validation Maturity
Summary and Recommendations
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 3: MPS-LO OVERVIEW
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 4: Space Environment In-Situ Suite (SEISS) Magnetospheric Particle Sensor - Low Energy (MPS-LO)
Primary purpose: measure GOES frame charging and spacecraft charging environment
Electrostatic analyzers (2 ESAs per species) with Microchannel plates (MCPs) as the detection elements
Measures electrons and ions within 30 eV – 30 keV over 15 logarithmically-spaced energy channels
14 angular zones (12 unique) in the north-south plane and 2 background zones
  15 deg FOV angular zones with combined total 180 deg FOV
  30 deg overlap in FOV across 4 angular zones should yield similar fluxes
L1b product: low energy magnetospheric electron and proton fluxes (counts/cm2-s-str-keV)
  2 x 14 x 15 = 420 flux time-series in MPS-LO L1b data
SEISS MPS-LO designed, built, tested and calibrated by Assurance Technology Corporation (ATC)
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 5: SEISS MPS-LO FOV
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 6: SEISS MPS-LO
An individual energy step observation interval is slightly less than 1/16 of a second and all 15 energy steps are observed during a 1 second cycle
A dead time correction is applied to each observation interval in accordance to the electronics event processing time
Ground processing uses the background zone observations to remove penetrating radiation from the particle observations
Magnetic field vector observations from the GOES magnetometer is essential for calculating pitch angle for each telescope / angular zone
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 7: A Few Comments About MPS-LO
NOAA’s first plasma instrument on GOES!
Measures thermal population (~1keV at geosynchronous)
Measuring the low energy component of a plasma is notoriously difficult – Plasma instruments are known to require special care.
Surface charging on spacecraft is associated with <50 keV electrons [NASA, “Mitigating in-space charging effects – a guideline,” NASA Technical Handbook NASA- HDBK-4002A, March 3, 2011].
Electrostatic discharge (ESD) due to spacecraft charging causes most of the environmentally related anomalies on spacecraft, and surface charging has caused the most serious anomalies, i.e., those that have resulted in the loss of the mission [Koons et al., 1999].
In addition to its operational use at SWPC, we are looking forward to making this data available to the science community.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 8: MPS-LO Primary Operational Use
SEISS MPS-LO L2 1m averages and the Density and Temperature Moments and Level of Spacecraft Charging algorithm are phase 2 algorithms. MPS-LO 5m averages is phase 3.
L2 densities and temperatures are also used in the L2 Magnetometer Magnetopause Crossing Detection algorithm as an indicator of magnetopause crossings at geosynchronous.
The L2 Moments algorithm identifies ion line signatures of SC charging and uses the ion line to determine the spacecraft potential:
When an ion line is not present, the L2 Moments routine uses an empirical relation between densities, temperatures, and level of charging to monitor SC charging.
1 sec.
Level 1b
Level 2
Data stream for NOAA-NWS SWPC
MPS-LO Density and Temperature Moments and Level of Spacecraft Charging
GOES-16 MPS-LO spectrogram of count rates, with ion line signatures of spacecraft charging. The first ion line enhancement appears near 1 keV at ~5:30-6:30 UT during an eclipse period.
UT Hours
On 22 March 2017
Backgrounds from heightened radiation belt fluxes
Ion line
MAG L1b 1m averages
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 9: Review of Beta/PLT Maturity
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 10: Review of Beta/PLT Maturity
GOES-18 SEISS achieved Beta status on August 1, 2022
ATC MPS-LO PLT Status
  PLT-SEI-005: MPS-LO Initial Turn-On of Electron and Ion Microchannel Plate High Voltage (MCP HV) – Final report delivered 2022-09-01
  PLT-SEI-004: MPS-LO Initial In-Flight Calibration (IFC) Execution – Final report delivered 2022-09-15
PLPT Overview
  PLPT-SEI-003: MPS-LO Zone Cross-Comparison
  PLPT-SEI-005: Cross Satellite Comparison of Trapped Particles
  PLPT-SEI-006: Backgrounds Trending
MPS-LO Specific ADRs/WRs
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 11: NCEI Beta Assessment GOES-18 MPS-LO 2022-07-19
Zones 1-7R Electrons
Zones 6L-12 Electrons
Zones 1-7R Ions
Zones 6L-12 Ions
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 12: NCEI Beta Assessment GOES-17 MPS-LO 2022-07-19
Zones 1-7R Electrons
Zones 6L-12 Electrons
Zones 1-7R Ions
Zones 6L-12 Ions
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 13: PLT-SEI-005: MPS-LO HV Activation and Bias Optimization
Determine optimal initial on-orbit HV bias voltage levels for electron and ion microchannel places (EMCP, IMCP)
  Preflight calibration: EMCP = step 24, IMCP = step 25
  EMCP and IMCP bias set to the sub-operating step 20 and increased incrementally to step 27 with data accumulation at each step
  IMCP bias increased to step 28
  EMCP and IMCP bias decremented to step 20 with data accumulation at each step
  Total test time period: 2022-07-01 11:00:17 – 13:16:22
Goal is to set HV low enough to avoid excessive noise and crosstalk while high enough to operate in/near saturation (uniformity of charge exiting MCPs with individual particle hit)
Conclusion: ATC recommend no changes to the MPS-LO MCP HV baseline settings (EMCP = step 24, IMCP = step 25)
Nominal retest cadence is six months
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Pass

## Slide 14: PLT-SEI-004: MPS-LO IFC Test
Measure the electron and ion zone thresholds at orbital extremes
  Expect thresholds to vary with temperature
  Ensure no difference in thresholds based on MCP bias state
IFC test completed for MPS-LO ion and electron MCP HV bias optimization
  IFC successfully executed with high voltage disabled and enabled
  IFC results are in-family with pre-launch results and current GOES operational spacecraft
  Temperature dependencies are comparable with ground test
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Pass

## Slide 15: Path to Provisional - PLPTs
Title | Activity | Data Needed | Success Criteria
A.2.2 MPS-LO Zone Cross-Comparison [PLPT-SEI-003] | Evaluate PADs, determine relative sensitivity/response among MPS-LO zones, and calculate inter-calibration factors for Z6L/Z6R and Z7L/Z7R for all 15 energies and for both electrons and ions. | One month of continuous data. | [Provisional] PADs match science expectations and zones agree within 25%.
A.2.4 Cross Satellite Comparison of Trapped Particles [PLPT-SEI-005] | Inter-calibrate MPS-LO with the same measurements on other operational GOES-R Series satellites. | One month of continuous data. | [Provisional] Differences quantified on data taken through validation phase.
A.2.5 Backgrounds Trending [PLPT-SEI-006] | Trend daily/weekly minima from MPS-LO ion and electron background zones | Two months of continuous data. | [Provisional] If dark noise can be observed above natural backgrounds, decreases with time.
MPS-LO Provisional PLPTs from RIMP v 2.0 (09 July 2021)
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 16: MPS-LO Specific ADRs/WRs
ADR | WR | Short Description | Status | Expected Closure
228 | 3961/7067 | SEISS record time mismatches with filenames | Closed | DO.07.01.00/DO.10.01.01
326 | 5424/5860 | MPS-LO GPA Changes | Closed | DO.07.01.00
449/ 596 | 5133/5800 | Make Arc Jet Flag available for SEISS MPS-LO | Closed | DO.07.00.00/DO.10.01.00
780 | 6845 | GPA update for SEISS MPS-HI and MPS-LO L1b IFC leakage | In-work | DO.12.00.00
1003 | 7506 | Flight delivery of SEISS FM3 LUTs | Closed | (2021/12/09)
1171 | 9326 | EXIS/MAG/SEISS ECEF_Z range | In-work | DO.12.00.00 (2023 spring)
1198/1199 |  | G16/G17 MPS-LO LUT updates | Closed | PR.09.08.08 (2022/10/02)
1209 |  | NCEI G18 SEISS LUT Actions (the real prelaunch LUTs) | Closed | PR.09.08.14 (2022/04/13)
1277 |  | G18 SEISS MPS-LO LUT Update | Closed | PR.08.08.38 (2022/11/09)
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 17: PRODUCT QUALITY EVALUATION
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 18: Summary of MPS-LO Anomalies (Common to all MPS-LO flight models to date)
High Backgrounds / Negative Fluxes
  Previous strategy with ADR 326 for G16 and G17 is also applied to G18
  Continued work on the empirical derivation of background removal coefficients
  Source of backgrounds is not fully understood
Interzonal Crosstalk
  G16/FM1 crosstalk seen in on-orbit data. Analysis to estimate correction coefficients not developed
  G17/FM2 crosstalk measured with ground/beam calibrations and is apparent in on-orbit data.
  G18/FM3 crosstalk observed on-orbit (slide 28)
Evidence for significant error in absolute values of reported fluxes (e.g., PLPT-SEI-003 and PLPT-SEI-005), which was also studied on G16 and G17 with no resolution. Possible sources of error in reported fluxes:
  Gradual decrease in in-band and background zone response with time (seen on slide 37)
  No crosstalk correction scaling of geometric factors
  May be insufficient or too much background removal
  To reduce crosstalk and noise, instrument not operated in full saturation
Degraded performance of lower energy ion channels due to persistent high backgrounds
  Previously studied on G16 and G17 with no resolution
Effect of arc jet firing
  Previously resolved on G16 and G17 with ADR 596 and resolution is true for G18
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 19: High Backgrounds / Negative Fluxes – ADR 326 and Empirical Derivation of Background Removal Coefficients at NCEI
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 20: G18 MPS-LO L1b Fluxes 2022/08/12 (Prior to implementation of ADR 1277 MPS-LO LUT Update)
Negative flux (black pixels) due to high backgrounds: Corrected counts = total counts – background counts
Addressed by ADR326/WR5860 MPS-Lo GPA Changes, in build DO.07.01.00, and ADR1277 (LUT update).
Right Sensor Head Ions
Black pixels indicate negative flux (non-physical), due to excessive background subtraction.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Left Sensor Head Ions

## Slide 21: MPS-LO Background Removal
Unilluminated
Background
Zone
Penetrating radiation / Backgrounds
BG
In-band
In-band
(electrons or ions)
~
Sweep voltage
(selects 15 energies)
14 zones
(viewing directions)
Simplified background (BG) removal expression :
Corrected Counts = Total counts – BG Removal Coefficient x BG Counts
Pre- ADR326 CDRL80 expression has 4 BG removal coefficients, one for each species and sensor head, all were set to one in the launch version of the MPS-LO LUT. Analysis shows need for energy and zone dependent BG removal coefficients with values <1 for most energy-zone pairs. Need to characterize in-band vs. BG response for 14 zones x 15 energies for electrons and ions.
Illuminated zone accumulates
both in-band and BG counts
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 22: Three Actions Addressing MPS-LO High Backgrounds / Negative Fluxes
ADR326: Implement energy dependent background subtraction, previously energy-sweep averaged (Also involves non-trivial modification to error term – Complete description in ADR_326_Addendum_Revised_10-2-18.pdf.)
ADR326: Generalization of background (BG) removal coefficient to include zone and energy dependence
  (2✕2) → (2✕2✕7✕15)
NCEI: Empirical derivation of BG removal coefficients (by demanding positive fluxes, averaged over some interval, using many months of data).
m = 1,2 for electrons and ions respectively.
k = 1,2 for right and left sensor heads.
r = 1-4 is species & sensor head group index,
for ions-R, ions-L, electrons-R, electrons-L.
i = 1-7 is zone index for 7 zones in each SH.
j = 1-15 is energy channel index.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 23: Empirical Determination of Background Removal Coefficients
The corrected counts M in a given zone and energy channel are the total counts Ntot. minus the counts from penetrating radiation
,
where i and j are the is the zone and energy indices, respectively, and the counts in zone i due to penetrating radiation is assumed to be proportional to the counts in the associated (same species and sensor head) background zone NBG.
If we demand that the flux averaged over some time-interval is non-negative, this places an upper bound on the background removal coefficient.
where the overbar denotes an average over some time interval (e.g., 5 min. averages from 1s data). An upper bound for can be determined empirically by taking the minimum value of attained over months of data
Set ,
where k is index for the averaging interval, and we take minimum from many such intervals. An upper bound of 1 is also placed on the C coefficients.
We only use cases
where and , so that rel. Poisson error associated with C is < 20%.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 24: Empirical Background Removal Coefficients
Ion Background Removal Coefficients
Electron Background Removal Coefficients
ADR 1277 G18 SEISS MPS-LO LUT Update implemented the following BRC tables.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 25: Verification of ADR 1277
Verification of ADR 1277: G18 L1b flux and error in agreement with NCEI’s in-house L0 to L1b result to within 1.E-7 (using single precision floats).
Right sensor head electrons
Left sensor head electrons
Flux error term for R & L SH electrons and ions
Rel. difference < 1.E-7 in all cases.
Verification of flux error term
(L1b – NCEI_in-house_L0_to_L1b) / NCEI_in-house_L0_to_L1b
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Left sensor head ions
Right sensor head ions
Verification of fluxes for ~13hrs of test data

## Slide 26: Background Removal Example 2022/08/10
Negative flux (black pixels) due to high backgrounds: Corrected counts = total counts – background counts
Addressed by ADR326/WR5860 MPS-Lo GPA Changes and ADR1277 (LUT update).
No Empirical Background Removal Coefficients
Black pixels indicate negative flux (non-physical), due to excessive background subtraction.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
New Empirical Background Removal Coefficients

## Slide 27: Additional MPS-LO Anomalies
Interzonal Crosstalk
  Partial resolution on slide 27 of G16/G17 Full Validation PS-PVR
Effect of High Backgrounds on Low Energy Ions Measurement
  Low signal to noise instrument/environment issue – No L1b resolution.
Effect of Arc Jet Firing
  Resolved with ADR596
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 28: Crosstalk Between Zones Seen in G18 MPS-LO Background Zones
Energy-sweep independent backgrounds from penetrating radiation
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Crosstalk is apparent in energy-sweep dependent count rates in electron background zones in on-orbit data. (Counts from penetrating radiation/particles should be independent of energy-sweep.)
Right sensor head
electron background zone
Left sensor head
electron background zone
Right sensor head
ion background zone
Left sensor head
ion background zone
G18 MPS-LO Backgrounds 2022/08/01
Energy-sweep dependent crosstalk from illuminated zones

## Slide 29: Effect of High Backgrounds on Low Energy Ions Measurement
High radiation belts
(high backgrounds)
Moderate geomagnetic storm day
(elevated ions)
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Energy independent count rate at lower energies dominated by background counts.
Signal to background ratio low on days when backgrounds are high.
Similar result for G16 MPS-LO
This is an instrument issue due to unexpectedly high background rates. There is no known resolution.

## Slide 30: MPS-LO Provisional Post Launch Product Tests (PLPTs)
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 31: MPS-LO PLPTs Supporting Provisional Maturity
PLPT-SEI-003 MPS-Lo Zone Cross-Comparison
PLPT-SEI-005 Cross satellite Comparison of Trapped Particles
PLPT-SEI-006 Backgrounds Trending
* PLPTs performed using NCEI in-house L0 → L1b with 2022/09/21 updated background removal coefficient table, which should now be in OE LUT. NCEI in-house L0 → L1b data verified in ADR 1277 to be within machine precision of the ground system generated L1b (slide 25).
** All three MPS-LO PLPTs given partial passes – Results from full validation continuations of these PLPTs will need to be applied to bring G18 MPS-LO into compliance with requirements.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 32: PLPT-SEI-003: MPS-LO Zone Cross-Comparison – Method (1/4)
Compares fluxes measured by pairs of MPS-LO angular zones when central pitch angles (α) are the same or supplementary.
Under such conditions, zones should observe the same flux at the same energy
Steps:
  Input 1-min MPS-LO fluxes and pitch angles calculated from MAG BRF
  Identify whenever a pair of zones has pitch angles within 1 degree and pitch angles varying more slowly than 30 deg/minute
  Use supplementary angles of pitch angles > 90 degrees to have sufficient matches (assumes up- and down-going trapped fluxes are the same on a 1-minute timescale)
  Report observed fluxes for each pitch angle match
  Estimate scaling factors optimally based on set of all matches
Ref: Rowland, W., and R. S. Weigel (2012), Intra-calibration of particle detectors on a three-axis stabilized geostationary platform, Space Weather, 10, S11002, doi:10.1029/2012SW000816.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Zone A look dir.
Zone B look dir.
B (magnetic field dir.)
α
α
Zone A look dir.
Zone B look dir.
B
α
α+90o
OR

## Slide 33: PLPT-SEI-003 – Method (2/4)
90°
60°
30°
0°
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Example: Ion Pitch Angles, August 2022 (supplementary angles of p.a. > 90 deg)
Dots indicate pitch angle matches
1,000-10,000 matches per month per energy-zone pair
Zone

## Slide 34: PLPT-SEI-003 – Method (3/4)
Use Nelder-Mead method to minimize the following objective function: sum of squares of the following difference term (~30k terms [matches] in the sum per energy)
To reduce effect of statistical noise, use only matches where both fluxes correspond to at least 1000 counts/minute
A non-linear problem: need a good initial guess for the scale factors in order for the solution to converge
  1000-6000 iterations needed, depending on the energy and species
  2.5 (4.3) hours needed for three months, 14 zones, 15 energies of ions (electrons) on macOS 2.4 GHz 8-Core Intel Core i9
Because they are only valid in a relative sense, scale factors must be normalized – here, normalized to those of Zone 8 (arbitrary choice)
Ref: Rowland, W., and R. S. Weigel (2012), Intra-calibration of particle detectors on a three-axis stabilized geostationary platform, Space Weather, 10, S11002, doi:10.1029/2012SW000816.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 35: PLPT-SEI-003 – Scale Factor Results (4/4)
Multiplicative scale factors plotted vs. energy channel for each zone
Marginal consistency month-to-month indicates need to do this on a monthly basis
Scaling of up to an order of magnitude needed for some zone-energy channels.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 36: PLPT-SEI-003 – Overlapping Zones Comparison (Z6R/Z6L and Z7R/Z7L)
Using 3 months of 1 min. averaged fluxes, the percent error with respect to the avg. of the overlapping pair is calculated for each 1m sample (i.e., we split the difference, thus errors associated with the 2 overlapping zones are the same). Then, the mean value of errors from 3 months of 1 min. averages is calculated.
The error computed by comparing fluxes from overlapping zones (left figure) is in close agreement with error computed from the scale factors obtained by minimizing error associated with cross comparisons of all zones described on previous slides (right figure), showing verification of results by an independent method.
Error computed from Z6 and Z7 scale factors
determined with method described on previous slides
3-month mean of error of each overlapping zone with respect to average of overlapping pair
Partial pass
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 37: PLPT-SEI-005: Cross Satellite Comparison of Trapped Particles
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
G16 EMCP changed from step 25 to 26 on 2019-08-06 13:56:00 UT.
G16 drift 89.5º to 75.2º west
G17 drift 89.5º to 137.2º west
G18 drift 89.5º to GOES west location
Comparison of G16, G17 & G18 MPS-LO Z10 E12 fluxes
Long term trending of day-averaged MPS-LO fluxes show steady non-geophysical decline.
Similar orders-of-magnitude decline is seen in all electron zone-energy channels.
Ion zone-energy channels show significantly less decline (420 analogous ion and electron flux plots created, available on request).
Steady decline in MPS-LO response since launch makes cross-comparison analysis ineffective.

## Slide 38: PLPT-SEI-005: Cross Satellite Comparison of Trapped Particles
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Note similar variations in G17 and G18 MPS-LO Z10 E12 fluxes, but initially rapid decline in G18 reported fluxes
Comparison of G16, G17 & G18 MPS-LO Z10 E12 fluxes

## Slide 39: PLPT-SEI-005: Cross Satellite Comparison of Trapped Particles: Summary
Previously: G16 & G17 MPS-LO Statistical Comparisons with Denton et al. [2015; 2016] model, based on 1990-2007 observations from the Magnetospheric Plasma Analyzer (MPA) instruments on-board seven Los Alamos National Laboratory (LANL) satellites at GEO (e.g., shown in G16/G17 Full Validation PS-PVR, slides 45 & 46).
Geophysically different eV-keV ion and electron flux levels at GOES east and west locations make direct comparisons between G16 and G17 MPS-LO units ineffective for determining relative responses.
Proximity of G17 and G18 warrant direct comparisons; however, decline in MCP gain in all MPS-LO units since launch (seen on previous slide) make G17 and G18 MPS-LO cross-comparison analysis ineffective.
Corrections for decline in MCP gain are under development at NCEI.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Partial pass

## Slide 40: PLPT-SEI-006: Backgrounds Trending
Extensive analysis of G18 MPS-LO backgrounds were performed for empirical determination of G18 MPS-LO BG removal coefficients.
Ratio of G18 MPS-LO in-band to background response characterized with determination of background removal coefficients.
As shown in previous MPS-LO PS-PVR, the source of backgrounds is not fully understood (mainly radiation belt electrons, but at lower energies than expected).
  Simulations show that penetration of MPS-LO box by electrons should be limited >=MeV energies.
  There is good correlation between MPS-LO backgrounds and 100s of keV electrons (400-800 keV).
In general, in-band and background zone responses have been gradually decreasing since launch (e.g., slide 37).
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.
Partial pass

## Slide 41: PROVISIONAL MATURITY ASSESSMENT
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 42: Summary of G18 MPS-LO Instrument/GPA Issues Identified and Resolution Status
High Backgrounds / Negative Fluxes
    Resolution: ADR 326 MPS-LO GPA Changes, ADR 1277 G18 MPS-LO LUT update, and empirical derivation of background removal coefficients by NCEI.
    Items above are completed, but BG removal coefficients will need additional updates.
Interzonal crosstalk
    Observed on-orbit with G16, G17, and now G18 MPS-LO.
    MPS-LO HV optimized to reduce crosstalk but a complete solution is not currently available.
Mitigating the effect of arc jet firing in operation – ADR 596
Evidence for significant error in absolute values of reported fluxes (e.g., PLPT-SEI-003: MPS-LO Zone Cross-Comparison and PLPT-SEI-005: Cross Satellite Comparison of Trapped Particles) – Geometric factors may need adjustment.
    MCP decline has a significant impact on absolute values of reported fluxes.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 43
GOES-18 MPS-LO
Performance Baseline Status
Assessment at Provisional
MRD ID | Description | MRD Requirement | MIT/LL Predicted Performance Baseline | NCEI Assessment at Provisional | Related PLPTs | Status
1987 | MPS-LO Orthogonality/ Coverage | 5 directions | 14 directions (zones) | 14 directions (zones) | -03 | PASS
1991 | MPS-LO Measurement Range | 30 eV – 30 keV (protons and electrons) | 30 eV – 30 keV | 0.025-30.8 keV (electrons) 0.030-30.0 keV/q (ions) [values from CDRL79] | -05 | PASS
1992 | MPS-LO Measurement Accuracy | 45% @min flux, 25% @10X min flux | 420 separate accuracies (2 species x 14 angles x 15 energies); Worst case for 10x min flux: Electrons: 23%; Ions: 20% Worst case for min flux: Electrons: 40%; Ions: 38% | Absolute accuracy cannot be verified on-orbit. Results of PLPT-03 and factor of ~5 decline in MCP response (in some zones) indicate that 25% relative accuracy not currently achieved on-orbit. | -03, -05, -06 | Results of PLPT-03, 05 and 06 will need to be applied
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 44: Provisional Validation
Preparation Activities | Assessment
Validation activities are ongoing and the general research community is now encouraged to participate. | Validation activities are ongoing. Results have been discussed with SWPC. Release of data by NCEI will enable research community participation.
Severe algorithm anomalies are identified and under analysis. Solutions to anomalies are in development and testing. | Continued analysis and updates of the NCEI empirical background removal coefficients.
Incremental product improvements may still be occurring. | Product improvements will result from steps to mitigate known (on previous chart) and any yet to be discovered anomalies.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 45: Provisional Validation
End State | Assessment
Product performance has been demonstrated through analysis of independent measurements obtained from select locations, periods, and associated ground truth or field campaign efforts. | G18/FM3/SN104 MPS-LO ground/beam calibrated. Performance tests included in provisional PLPT results
Product analysis is sufficient to communicate product performance to users relative to expectations (Performance Baseline). | Yes, to the limited extent that absolute accuracy can verified on-orbit.
Documentation of product performance exists that includes recommended remediation strategies for all anomalies and weaknesses. Any algorithm changes associated with severe anomalies have been documented, implemented, tested, and shared with the user community. | A complete description of known anomalies will be included in readme file with release of data to public.
Testing has been fully documented. | In this presentation and supplemental materials.
Product is ready for operational use and for use in comprehensive cal/val activities and product optimization. | Readiness for operational use at provisional PS-PVR review.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 46: PATH TO FULL VALIDATION MATURITY
GOES-18 MPS-LO L1b Product Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 47: Path to Full Validation
PLPTs to be conducted prior to Full Validation.
Risks to getting to Full, both from L1b product definition and user (SWPC) feedback
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 48: Path to Full Validation - PLPTs
Title | Activity | Data Needed | Success Criteria
A.3.3 MPS-LO Zone Cross-Comparison [PLPT-SEI-003] | Evaluate PADs, determine relative sensitivity/response among MPS-LO zones, and calculate inter-calibration factors for Z6L/Z6R and Z7L/Z7R for all 15 energies and for both electrons and ions. | Three months of continuous data. | Zones agree to within 25% absent spacecraft charging similar to or greater than channel energy.
A.3.5 Cross Satellite Comparison of Trapped Particles [PLPT-SEI-005] | Inter-calibrate MPS-LO with the same measurements on other operational GOES-R Series satellites. | Three months of continuous data. | Differences quantified on data taken through validation phase.
A.3.6 Backgrounds Trending [PLPT-SEI-006] | Trend daily/weekly minima from MPS-LO ion and electron background zones | Four months of continuous data. | If dark noise can be observed above natural backgrounds, decreases with time.
MPS-LO Provisional PLPTs from RIMP v 2.0 (09 July 2021)
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 49: Path to Full Validation – Risks
Risk Name | Description | Impact | Mitigation and Schedule
High Backgrounds / Negative Fluxes | Higher than expected backgrounds results negative (background subtracted) fluxes in L1b. | Degrades accuracy of reported fluxes. | Addressed by ADR 326 MPS-LO GPA Changes, ADR 1277 G18 MPS-LO LUT Update, and empirical derivation of background removal coefficients. Background removal coefficients will need to be periodically updated.
G18/FM3 Crosstalk | Interzonal crosstalk observed in background count rates on-orbit. | May introduce up to 10% error in fluxes (severity dependent on HV setting). | G18 MPS-LO HV optimized to reduce crosstalk but a complete solution is not currently available.
Significant error in absolute values of reported fluxes. | Evidence for significant error in absolute values of reported fluxes from Provisional PLPT results. | Reported fluxes up to an order of magnitude in error - Severely degrades operational and science quality of data. | Full validation PLPTs PLPT-SEI-003 MPS-LO Zone Cross-Comparison and PLPT-SEI-005 Cross satellite Comparison of Trapped Particles, and full assessment of changes in MCP response will provide necessary scaling of geometric factors to be included in future LUT updates.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 50: SUMMARY & RECOMMENDATIONS
GOES-18 MPS-LO L1b Products Provisional PS-PVR
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 51: Summary
The G18/FM3 MPS-LO anomalies identified that negatively impact performance are
  High Backgrounds / Negative Fluxes – Addressed by ADR 326 MPS-LO GPA Changes, ADR 1277 G18 MPS-LO LUT Update, and empirical derivation of background removal coefficients by NCEI. Background removal coefficients will need additional updates.
  Interzonal Crosstalk – Observed in energy-dependent background count rates on-orbit.
  Evidence for significant error in absolute values of reported fluxes (e.g., PLPT-SEI-003: MPS-LO Zone Cross-Comparison and PLPT-SEI-005 Cross satellite Comparison of Trapped Particles) – Geometric factors will need periodic adjustment.
MPS-LO uses the ion-line to identify SC charging and quantify the level of charging. If an ion line is not present, MPS-LO L2 processing uses an empirical relation between densities, temperatures, and level of charging to monitor SC charging.
Densities and temperatures from L2 processing will be affected by error in absolute values of reported fluxes, but the empirical relation used to determine level of charging from density and temperature (yet to be determined for MPS-LO) will compensate for errors in absolute value of reported fluxes.
G18/FM3 MPS-LO is currently performing as needed to fulfill an intended operational use, i.e., identifying SC charging and level of charging.
A path to achieving accurate fluxes and moments has been identified (on following slide).
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 52: Recommendations
NCEI recommends G18 MPS-LO L1b data for transition to provisional status.
    Continue performing MPS-LO HV optimization procedure every 6 months. (Initial ATC recommendation to perform HV optimization every 6 months – May modify to once per year as MCPs stabilize.)
Monitor effect of changes in HV setting on observed count rates.
Background removal coefficients will need to be updated regularly. We recommend performing empirical determination of background removal coefficients and MPS-LO LUT update following each HV optimization.
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 53: Recommendations Cont.
Scale geometric factors using results from [Full] PLPT-SEI-003 MPS-LO Zone Cross-Comparison. This will bring MPS-LO into agreement with itself. The Zone Cross-Comparison procedure should be repeated periodically to check for self-consistent agreement among MPS-LO zones and scale geometric factors as needed.
Perform careful comparisons with current data from other missions (e.g., LANL MPA), to insure quality of data for science community and support space weather modeling and prediction at SWPC. This activity is included in [Full] PLPT-SEI-005. This will provide scaling of geometric factors needed to correct error in absolute values of reported fluxes.
Empirically determined parameters used in the SEISS Level 2 Density and Temperature Moments and Level of Spacecraft Charging algorithm need to be tailored to MPS-LO.
MPS-LO is NOAA’s first plasma instrument on GOES –
We need to take special care!
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 54: Backup Slides
PLPT-SEI-005: Cross satellite Comparison of Trapped Particles
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 55: In Development Correction of Decline in MCP Gain for NCEI Retrospective/Science
High Backgrounds / Negative Fluxes
  Previous strategy with ADR 326 for G16 and G17 is also applied to G18
  Continued work on the empirical derivation of background removal coefficients
  Source of backgrounds is not fully understood
Interzonal Crosstalk
  G16/FM1 crosstalk seen in on-orbit data. Analysis to estimate correction coefficients not developed
  G17/FM2 crosstalk measured with ground/beam calibrations and is apparent in on-orbit data.
  G18/FM3 crosstalk observed on-orbit (slide 28)
Evidence for significant error in absolute values of reported fluxes (e.g., PLPT-SEI-003 and PLPT-SEI-005), which was also studied on G16 and G17 with no resolution. Possible sources of error in reported fluxes:
  Gradual decrease in in-band and background zone response with time (seen on slide 37)
  No crosstalk correction scaling of geometric factors
  May be insufficient or too much background removal
  To reduce crosstalk and noise, instrument not operated in full saturation
Degraded performance of lower energy ion channels due to persistent high backgrounds
  Previously studied on G16 and G17 with no resolution
Effect of arc jet firing
  Previously resolved on G16 and G17 with ADR 596 and resolution is true for G18
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 56: Correction in-work for MPS-LO gain decline for NCEI retrospective/science processing
Analysis to determine corrections for each zone-energy channel:
  Nonlinear fit (black trace) to: (3 parameters)
  Dead-time corrected day-averaged count rate scaled up to fit at first light, i.e., Correction factor = fit(t=0) / fit(t)
    Assumes relatively steady long-term rates over several year period. (Figure on following slide shows this is a reasonable assumption at higher MPS-LO energies.)
    Does not account for different MCP HV settings on-orbit and during ground/beam cal.
    Assumes relatively steady long-term rates over several year period. (Figure on following slide shows this is a reasonable assumption at higher MPS-LO energies.)
    Does not account for different MCP HV settings on-orbit and during ground/beam cal.
  Must recompute BG removal coefficients after applying corrections.
  Some possible ways to constrain/modify fits include:
    Demand agreement between G16, G17 and G18 MPS-LO long-term flux variations
    Demand similar G16, G17 and G18 MPS-LO fits in like zone-energy channels
    Choose fits that force G17 and G18 fluxes to agree, while both are operating at GOES west location
    Reproduce long-term MPS-HI flux variations at highest MPS-LO energies
G16 MPS-LO Electron Zone 7R, E15
Two fits performed then forced to agree on 2019-08-06.
Corrected
Uncorrected
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.

## Slide 57: Compare G16 MPS-LO 30 keV and G16 MPS-HI 72 keV
(Last day in plot is 2021-12-31)
Mean value
These GOES-18 data are preliminary, non-operational data and are undergoing testing. Users bear all responsibility for inspecting the data prior to use and for the manner in which the data are utilized.